## Slide 1
The System Function
Notes: Here we will examine the system function and its experimental determination.

## Slide 2
Learning Objectives
Combination of transfer functions and Thevenin equivalent
pulser source term into a single factor – the system function
Relationship between the system function and the system
"efficiency" factor
Experimental determination of the system function
by deconvolution
Examples of synthesizing an entire ultrasonic measurement system
and predicting the measured signals
Notes: One can collect the sound generation and sound reception transfer functions together with the Thevenin equivalent voltage source of the pulser into a single function called the system function which we can obtain experimentally through models and measurements . We will relate this system function to a closely related function called the system efficiency factor which has been used in earlier studies. We will give a number of examples of the use of a system function to predict the detailed characteristics of measured ultrasonic signals.

## Slide 3
The System Function
acoustic/elastic
processes
Notes: Here we see a complete ultrasonic NDE system characterized by three transfer functions and a “reduced” model of the system where the sound generation and reception transfer functions are combined with the Thevenin equivalent voltage source to model the system as simply the acoustic/elastic transfer function where the system function serves as the driving input.

## Slide 4
… system function
The System Function
There is a relationship between s(w) and the system “efficiency”
factor, b(w), [ Schmerr, Fundamentals of Ultrasonic Nondestructive
Evaluation, 2nd Ed. ] :
piston velocity
average incident pressure
Notes: In earlier studies the acoustic/elastic transfer function was defined as the average incident pressure acting on the receiving transducer divided by the density and wave speed multiplied by the uniform velocity on the face of the transmitting transducer, as shown. When we relate the received voltage to this acoustic/elastic transfer function and a driving input, the driving input is called the system efficiency factor, which different from the system function. We can use either the system function or the system efficiency factor but we must be aware of the differences.

## Slide 5
where
For piston transducers,
high frequency
The System Function
SR , ST … areas of receiving and transmitting transducers
Notes: If we compare the system function and system efficiency factor and the corresponding acoustic/elastic transfer functions then for piston transducers and at high frequencies, where the acoustic radiation impedance is just a plane wave acoustic impedance, we see the system function and system efficiency are just proportional to each other.

## Slide 6
The System Function
Two methods to find s(w):
A. Measure all the components needed to obtain tR(w), tG(w),
Vi(w) by the methods indicated previously. Then combine
these measurements to give
Notes: Now, consider how we can experimentally measure the system function. There are two methods. The first is to relate the system function to all of the underlying components and measure those components. This is lengthy but possible since we have demonstrated how to measure or model all of those components.

## Slide 7
The System Function
deconvolution by
straight division
Wiener filter
deconvolution
B. Measure VR(w), model tA(w), compute
Notes: The second method, which is much simpler, is to measure the received voltage in a calibration setup where the acoustic/elastic transfer function can be modeled explicitly and then obtain the system function by deconvolution (division in the frequency domain). Shown is an example pitch-catch setup where the two transducers are placed in an aligned fashion in a water bath and the acoustic/elastic transfer function is given. To stabilize the deconvolution it is not implemented by straight division but through the use of a Wiener filter that includes a small filter constant, .

## Slide 8
The System Function
Comparison of system factor obtained with both methods.
method A : synthesized
method B: measured
0
2
4
6
8
10
12
14
16
18
20
0
2
4
6
8
x 10
-2
Frequency (MHz)
synthesized
measured
s( f )
volts/MHz
Notes: Here is an example of using both of these methods to obtain the system function. We see that generally we have good agreement.

## Slide 9
(a)
(b)
(c)
(d)
(e)
(f)
The System Function
Example setups where tA is known
Notes: Here are a wider class of setups where the acoustic/elastic transfer function can be modeled explicitly and used to determine the system function through deconvolution. Most of these are for pulse-echo setups, but we do have one setup suitable for a pitch-catch or through-transmission inspection.

## Slide 10
The System Function
If all the electrical and electromechanical components of an ultrasonic measurement system are measured to obtain s(w), then this s(w) in conjunction with a knowledge of tA can used to determine the output voltage of a measurement system.
The following examples are shown for this setup:
pulser
receiver
Notes: Although direct measurement of the system function through deconvolution is much simpler than measuring and combining all the underlying system components , a knowledge of those components lets us examine how the measured voltage is affected by all the individual components so that we could, for example, predict how a change of a transducer(s) sensitivity affects the measured voltage when designing a measurement system.
Here, we will give a number of examples of the voltage versus time waveforms predicted for the setup shown compare to the actual measured signals when all the components of the system function are measured.

## Slide 11
The System Function
2.2
2.4
2.6
2.8
3
3.2
3.4
-0.6
-0.4
-0.2
0
0.2
0.4
0.6
time, msec
voltage, volts
Results for a pair of 5MHz, 6.35 mm diameter transducers.
solid line – directly measured output voltage in an ultrasonic pitch-catch setup.
dashed-dotted line – voltage signal synthesized by modeling all the ultrasonic
components and performing an inverse FFT to obtain the time domain signal.
Notes: First example. All the system components were measured and combined with the acoustic/elastic transfer function to obtain the measured voltage in the frequency domain that was then inverted into the time domain and compared with the directly measured voltage signal.

## Slide 12
The System Function
2
2.5
3
3.5
4
4.5
-0.4
-0.3
-0.2
-0.1
0
0.1
0.2
0.3
0.4
time, msec
voltage, volts
Results for a pair of 2.2 5MHz, 12.7 mm diameter transducers.
solid line – directly measured output voltage in an ultrasonic pitch-catch setup.
dashed-dotted line – voltage signal synthesized by modeling all the ultrasonic
components and performing an inverse FFT to obtain the time domain signal.
Notes: Second example. Again, all the system components were individually measured.

## Slide 13
The System Function
2.3
2.4
2.5
2.6
2.7
2.8
2.9
3
3.1
-0.4
-0.3
-0.2
-0.1
0
0.1
0.2
0.3
0.4
0.5
0.6
time, msec
voltage, volts
Results for a pair of 10 MHz, 6.35 mm diameter transducers.
solid line – directly measured output voltage in an ultrasonic pitch-catch setup.
dashed-dotted line – voltage signal synthesized by modeling all the ultrasonic
components and performing an inverse FFT to obtain the time domain signal.
Notes: Third example. Again, all the system components were individually measured.

## Slide 14
The System Function
voltage, volts
time, msec
2.4
2.6
2.8
3
3.2
-0.2
-0.1
0
0.05
0.15
voltage, volts
2.5
2.6
2.7
2.8
2.9
3
3.1
-0.15
-0.1
-0.05
0
0.05
0.1
time, msec
Results for a 5MHz, 6.35 mm diameter transducer.
Results for a 10 MHz, 6.35 mm diameter transducer
solid line – directly measured output voltage in the ultrasonic pulse-echo setup.
dashed-dotted line – voltage signal synthesized by modeling all the ultrasonic
components and performing an inverse FFT to obtain the time domain signal.
(a)
(b)
Notes: All the previous examples were for a pitch-catch setup. Here are examples for several transducers in the pulse-echo setup shown above.

## Slide 15
Homework problem
7.1
Notes: Homework problem from Chapter 7.