## Slide 1
The Ultrasound Reception Process
Notes: Here, we will describe all the elements of the reception process where the acoustic/elastic waves are transformed into a measured output voltage.

## Slide 2
Learning Objectives
Blocked force as a "lumped" source on reception
Thevenin equivalent transducer/sources on reception
Measurement of receiving cabling properties
Characterization of receiver as impedance, gain
and measurement of these receiver parameters
Reception transfer function
Ultrasonic system model
Notes: We will examine why the blocked force is important in the reception process and examine how this force acts as part of a Thevenin equivalent model of the receiving transducer and its acoustic sources. Again, we will examine the cabling and discuss how to model the receiver. All of these elements will then be combined into a complete model of the reception process.

## Slide 3
Transducer on reception
F
v
V
I
[TB]
Note: assumed directions
of v,I are now reversed from sound generation case so:
acoustic
sources ?
incident waves
scattered waves
receiving transducer
F , v
Notes: Like the sending transducer the receiving transducer can be modeled as a two port system but where now the acoustic port is the driving input port and the electrical is the output. Also note that the directions of the velocity and current are now changed from the sending case. These changes means that the elements of the transfer matrix used to characterize the transducer as a transmitter are in different locations in the transfer matrix when the transducer is used as a receiver, as shown. To make this model useful we must describe how the acoustic waves acting on the receiver can be defined in terms of a lumped source parameter.

## Slide 4
Transducer on reception
pinc
pscatt
v (x,w) n
F, force on
transducer
=
pinc
pscatt
FB , blocked
force
B
+
S
.
v n = 0
.
v (x,w) n
.
original problem
problem I
problem II
pscatt
FS , force due to motion of
transducer face
Notes: Consider a problem where a set of waves are incident on a receiving transducer, producing some scattered waves as well as some motion of the face of the transducer. We can consider this problem as the superposition of problem I, where the face of the transducer is held fixed, and problem II where the incident waves are absent and the transducer has the motion of the face of the original problem. The force on the transducer face for problem I is the blocked force, FB, while the force due to the motion of the face in problem II we will call Fs. Obviously the total force is the sum of these two forces.

## Slide 5
Fs
S
v (x,w) n
.
problem II : transducer is acting as a transmitter
pscatt
FB(w)
Zr (w)
a
F
v
[TB]
V
I
Transducer on reception
(1)
Notes: In problem two the waves generated are due only to the motion of the transducer face so this is identical to the case where the transducer is used as a transmitter except now the direction of the velocity is different. Thus, a minus sign appear when we relate the force and the velocity through the acoustic radiation impedance. The total force is then given by Eq. (1) and this relationship can be modeled as shown in the final figure where the blocked force acts as a source term in series with the acoustic radiation impedance at the input port. Thus, the blocked force is the natural force to calculate at the receiving transducer.

## Slide 6
FB = 2 Finc = twice the force due to the
incident waves only
B
v n =0
.
problem I
pscatt
2 Finc(w)
Zr (w)
a
F
v
[TB]
V
I
pinc
Note: many authors assume the incident and scattered
waves can be treated as plane waves incident on a fixed planar surface. In this case, as we showed previously:
Transducer on reception
Notes: As we showed previously, if we model the incident and scattered waves at the receiver locally as plane waves, then the blocked force is just twice the force in the incident waves when the receiving transducer is absent so we can replace the blocked force by twice this incident force in our model.

## Slide 7
Transducer on reception
Thevenin equivalent
reduced transducer and acoustic sources model
F
v
V
I
[TB]
V
I
Notes: We can replace the acoustic sources and transducer by a Thevenin equivalent system consisting of only a voltage source and an electrical impedance.

## Slide 8
reduced transducer and acoustic sources model
V
I
Transducer on reception
under open
circuit conditions
but, recall
so
open-circuit voltage
Notes: We can show that the open circuit voltage is just the blocked force times the sensitivity factor we discussed previously when the transducer is used as a transmitter.

## Slide 9
reduced transducer and acoustic sources model
V
I
Transducer on reception
To get the impedance, remove the blocked force source and
measure :
F
v
V
I
[TB]
but this is the same configuration as during
sound generation so
Notes: We can get the Thevenin equivalent impedance by removing the blocked force source, but this just yields the same configuration as when the transducer is being used as a transmitter so the electrical impedance here is the same as the one calculated when the transducer is being used as a transmitter.

## Slide 10
Transducer on reception
reduced transducer and acoustic sources model
V
I
Notes: Thus, the final Thevenin equivalent circuit is as given, showing that the transducer impedance and sensitivity characterize the transducer completely when used as both a transmitter or receiver.

## Slide 11
V1
I1
V2
I2
Receiving cable – transfer matrix model
[R]
V1
I1
V2
I2
R11 = R22 so changing directions on
currents does not change [R]
transmission line model
Notes: For the receiving cable we can again use the same two port system transfer matrix whose elements can be obtained with electrical measurements as shown previously.

## Slide 12
Receiving cable - experimental values
Cable - 5 ft
Fixture - 2.5 ft
Receiver
In immersion setups,
cabling and fixtures
holding the transducer
need to be characterized
0
5
10
15
20
0
0.5
1
1.5
0
5
10
15
20
-200
-100
0
100
0
5
10
15
20
0
50
100
0
5
10
15
20
-150
-100
-50
0
5
10
15
20
0
0.01
0.02
0
5
10
15
20
-110
-100
-90
-80
0
5
10
15
20
0
0.5
1
0
5
10
15
20
-200
-100
0
100
Frequency (MHz)
Frequency (MHz)
Amplitude
Phase
(Deg)
(Deg)
(Deg)
(Deg)
R22
R21 (1/W)
R12 (W)
R11
Notes: In immersion measurements there may be internal cables in fixtures holding the transducer as well as coaxial cables to the receiver but we can measure the transfer matrix of such combinations directly in the same manner, as shown here. The behavior of the results are in general agreement with what we expect from the transmission line model.

## Slide 13
Receiver - equivalent circuit
V0
I0
Z0 (w)
e
VR = K(w) V0
Z0 (w) … receiver impedance
e
K(w) … receiver gain
receiver
output voltage
Notes: The receiver part of the pulser will have some characteristic input electrical impedance and will amplify the signal. The amplification we can characterize by a frequency dependent gain factor. We will not include any low or high pass filtering in our receiver model since in quantitative NDE measurements we normally do not enable those filters. Such filters, however, can easily be added to our model if desired.

## Slide 14
Z0 (w) = V0 (w) / I0 (w)
K(w) = VR(w) / V0(w)
Receiver - experimental
transducer
cable
amplifier
VR
I0
V0
current probe
e
Notes: If we measure the input current and voltage as a function of frequency at the receiver input and the receiver voltage output we can calculate both the impedance and the gain.

## Slide 15
Ultrasonic System Components
current probe
Notes: The voltage is measured by simply tapping the voltage at a T-adapter while the current is measured by stripping off a small section of cable and clamping a commercially available current probe onto the exposed inner conductor. This assembly is placed in a small box as shown.

## Slide 16
Receiver - experimental
0
0.5
1
1.5
2
2.5
3
3.5
4
8
9
10
11
12
13
14
15
Amplitude
Frequency (MHz)
-40
-20
0
20
40
60
80
100
120
140
Phase (Deg)
amplitude
phase
Receiver gain
gain: 20 dB
attenuation: 0 dB
measured with
a 2.25 MHz
transducer
Receiver settings:
Notes: Shown are some measurements of the amplitude and phase of the receiver gain function at a particular gain and attenuation setting. The electrical inputs were generated in an ultrasonic setup where a 2.25 MHz transducer was connected to the receiver. Thus, the measurements here are only good for the bandwidth of this transducer. If we need to have a wider response we could use a different wide-band source or a different transducer.

## Slide 17
Receiver - experimental
Receiver impedance
0
0.5
1
1.5
2
2.5
3
3.5
4
0
200
400
600
800
1000
1200
-150
-100
-50
0
50
100
Frequency (MHz)
Amplitude (W)
Phase (Deg)
amplitude
phase
Notes: Here is a corresponding impedance measurement of the receiver.

## Slide 18
Reception process model
transducer +
sources
cabling
receiver
Notes: We can combine all the elements of the reception process into a single transfer function for reception as shown.

## Slide 19
Generation and Reception process models
Both the generation and reception process transfer functions
depend only on the transducer sensitivities and their electrical impedances:
Notes: Here we show the complete transfer functions for both the sending and receiving parts of an ultrasonic measurement system for a pitch-catch immersion setup. We see that the transducers impedance and sensitivity are the only parameters we need to completely characterize the role of the transducers in a measurement so we do not need to know their transfer matrix components explicitly.

## Slide 20
Ultrasonic System Model
acoustic/elastic
processes
Notes: We now have complete models of both the sound generation and sound reception parts of an ultrasonic measurement system, leading to the three LTI systems shown. If we combine the generation and reception process transfer functions with the Thevenin equivalent voltage source into a system function then we have an even simpler model consisting of a single LTI system where the system function acts as the input and the transfer function is the acoustic/elastic transfer function. Thus, if we have a practical way to measure the system function and if we can model the acoustic/elastic transfer function, we can predict the measured output voltage.

## Slide 21
Homework problem
5.1
Notes: Homework problem from Chapter 5.