## Slide 1
The Fourier Transform
Notes: The following slides will give a brief introduction to the Fourier transform. It is this transform that allows us to characterize completely an entire ultrasonic system in terms of its components in a direct and simple manner.

## Slide 2
Learning Objectives
definition of Fourier Transform and its inverse
reasons for use of the Fourier Transform
brief introduction to Fourier transform properties
delta function and its spectrum
definition of dB scale
definition of -6dB bandwidth
bandwidth - time domain connection
Notes: We will define the Fourier transform and its inverse and give the reasons why this transform is so useful. We will also describe some of the characteristics of Fourier transforms and give an example of a very important special case – the Fourier transform of a delta function.
We will define the decibel scale often used for ultrasonic signals and the characterization of the frequency domain spectrum of a signal in terms of its -6 dB bandwidth.
Finally, we will demonstrate the connection between the characteristics of a signal in the time domain and in the frequency domain.

## Slide 3
Fourier Transforms for UT Systems
vi(t)
vR(t)
Time domain Frequency domain
input voltage
output voltage
Fourier transform
inverse Fourier transform
Notes: Even though an ultrasonic measurement system inherently involves time-dependent signals such as the input voltage of the pulser or the measured output voltage at the receiver, as shown, it is easier, as stated earlier, to characterize all the elements of the measurement system in the frequency domain. Fortunately, the Fourier transform allows us to generate the frequency domain signals from those in the time domain. Similarly, the inverse Fourier transform allows us to reverse this process and generate the time domain signals from those in the frequency domain.

## Slide 4
… dimensions are volts-sec or volts-msec
or
for NDE problems t is usually in msec, f in MHz
Equivalent forms
Notes: Shown are the definitions we will use of the Fourier transform and its inverse. Those transforms can be written in terms of the frequency in rad/sec or in terms of cycles/sec (Hz). In NDE tests the time signals are often very short so they are typically measured in microseconds (μsec) and the frequencies are very high so that the corresponding frequency domain signals are typically measured in terms of millions of cycles/sec (MHz).
Note that if v(t) is a time domain voltage signal with t measured in seconds, its frequency spectrum V(f) has dimensions of volts-sec, but because we will often work entirely in the frequency domain, we will ignore this difference and refer to the frequency spectrum signals as simply volts. Other quantities such as pressure, velocity, current, etc. will follow this naming convention.

## Slide 5
A few properties of Fourier Transforms
If
then
Fourier Transforms
Notes: Because of the complex exponential term in the Fourier transform, the frequency domain signal V(f) will usually be complex even though the time domain signal v(t) is real.
Using this definition it is easy to derive a number of relationships between time domain signals and frequency domain signals. Two important cases are shown here where we see that a shifting (delay) of a time signal causes the frequency spectrum to be multiplied by a complex exponential term. Similarly, differentiating a time domain signal is equivalent to multiplication by an imaginary frequency dependent term in the frequency domain.

## Slide 6
t0
t
A
v(t)
Example Fourier transform
Notes: As a simple example of a Fourier transform, consider the box function shown in the time domain. In this case the Fourier transform is easy to calculate and we see that indeed this the transform is a complex function in the frequency domain.

## Slide 7
>> u=linspace(0, 5, 100);
>> u = u + eps*(u = = 0);
>> yf = exp(i*pi*u).*sin(pi*u)./(pi*u);
>> plot(u, abs(yf))
>> xlabel('f*t_0')
>> ylabel('abs(V/At_0)')
0
0.5
1
1.5
2
2.5
3
3.5
4
4.5
5
0
0.1
0.2
0.3
0.4
0.5
0.6
0.7
0.8
0.9
1
f*t
0
abs(V/At
0
)
Notes: If we evaluate the Fourier transform of the box function and plot its normalized magnitude versus non-dimensional frequency for positive frequencies we see most of the frequency content is in a low frequency “lobe” with decreasing content in side lobes as the frequency increases.

## Slide 8
>> plot(u, angle(yf))
0
0.5
1
1.5
2
2.5
3
3.5
4
4.5
5
0
0.5
1
1.5
2
2.5
3
3.5
Notes: If we plot the phase of the function we see a general linear increasing behavior with jumps of pi radians. The linear behavior comes from the complex exponential term in the function and the jumps of pi radians occur because of the periodic changes in sign of the sine function ( recall exp( i π) = -1 ).

## Slide 9
Concept of a delta function (impulse)
t0
t
A
v(t)
Consider taking the limit
with
1.0
t
f
Notes: If we imagine taking the limit of the box function as its time duration goes to zero but where we keep the product of its amplitude and duration at unity, then conceptually we obtain an infinite spike of zero duration at time t = 0, which we call a delta function. In the frequency domain the spectrum simply goes to a constant value of one at all frequencies. Thus, a delta function excites all frequencies equally which makes it an ideal input function to characterize a system’s frequency domain response.
We will indicate that the delta function is a key concept by placing a light bulb next to it. In later slides we will use the light bulb symbol to highlight other key concepts.

## Slide 10
Some properties of delta functions
t
t
1.0
t
t
unit step function
0.5
Notes: Here we list several important properties of delta functions. We see that a delta function multiplied by an “ordinary” function in an integral has the property of sampling that ordinary function at the location of the delta function. Similarly, we see that an integral of a delta function generates a unit step function at the location of the delta function.

## Slide 11
produces a Gaussian spectrum with a center frequency fc and
bandwidth bw where
frequency
V( f )
Notes: This slide shows an example time domain signal that generates a spectrum consisting of two real Gaussians, one centered at a positive frequency and one centered at a negative frequency. Only the Gaussian centered at the positive frequency is shown in the plot.

## Slide 12
bw = -6 dB bandwidth
frequency
bw
fc = center frequency
definition of a
ratio in decibels
a ratio ½ is
equivalent to -6 dB
Notes: Gaussians are often used to discuss signal characteristics. For example, we can define the bandwidth of the Gaussian as the width of the function, bw, where the function drops to one half of its maximum value. The value of ½, as measured in decibels (dB) is -6 dB so this is also called the -6 dB bandwidth. Similarly, we can call the frequency at the location of the maximum of the Gaussian as the center frequency of the signal.

## Slide 13
In real ultrasonic systems the spectrum is not symmetric, so how do we define a bandwidth and center frequency?
Vmax
Vmax /2
f1
f2
center frequency:
-6 dB bandwidth (in % of center frequency):
fc
Notes: In a real ultrasonic signal (whose spectrum is not symmetric like the Gaussian) we can determine the maximum magnitude of the spectrum and then find the frequencies where the signal is one half of that maximum value. Those two frequencies then can be used to determine both a center frequency and -6dB bandwidth for the signal as shown.

## Slide 14
- 6 dB bandwidth in % of center frequency
Transducer specifications (from the manufacturer)
Notes: Here is an example of a specification sheet obtained from a transducer manufacturer that shows plots of the time domain and frequency domain outputs of the transducer and where we see they give both the center frequency and -6 dB bandwidth.

## Slide 15
V( f )
[note scale differences]
v(t)
bw = 4 MHz
bw = 2 MHz
bw = 1 MHz
frequency, MHz
time, m sec
fc = 10 MHz
-20
-10
0
10
20
0
0.1
0.2
-1
-0.5
0
0.5
1
-1
0
1
-20
-10
0
10
20
0
0.2
0.4
-1
-0.5
0
0.5
1
-1
0
1
-20
-10
0
10
20
0
0.5
1
-1
-0.5
0
0.5
1
-1
0
1
Notes: We can use the time domain signal discussed previously that generates Gaussians in the frequency domain to examine how changes in the time domain and frequency domain are related. Here, we keep the center frequency fixed at 10 MHz and change the bandwidth. We plot both positive and negative frequencies to show that there are actually two Gaussians present. At the largest bandwidth we see that the time domain signal is the shortest, becoming wider in the time domain as the bandwidth becomes smaller. Note that the amplitude of the frequency domain signals also gets larger as the bandwidth decreases even though the amplitude of the time domain signal is unchanged. This behavior may not be as directly evident since we see the scale of the response in the frequency domain is changed for the various cases to make the plots more readable.

## Slide 16
% Gaussian_script
f= linspace(-20, 20, 200);
t =linspace(-1,1, 500);
subplot(3,2,1)
bw =4;
fc = 10;
[y,z]= Gauss_funcs(f,t,fc, bw);
plot(f, y)
subplot(3, 2, 2)
plot(t,z)
subplot(3, 2, 3)
bw =2;
fc = 10;
[y,z]= Gauss_funcs(f,t,fc, bw);
plot(f,y)
subplot(3,2,4)
plot(t,z)
subplot(3, 2, 5)
bw =1;
fc = 10;
[y,z]= Gauss_funcs(f,t,fc, bw);
plot(f,y)
subplot(3,2,6)
plot(t,z)
function [y, z] =Gauss_funcs(f,t,fc,bw)
a = sqrt(log(2))/(pi*bw);
y = sqrt(pi)*a*(exp(-(2*a*pi*(f - fc)).^2) + exp(-(2*a*pi*(f + fc)).^2));
z = cos(2*pi*fc*t).*exp(-(1/(4*a^2))*t.^2);
Notes: Here is the MATLAB code that generates the plots seen on the previous slide. The Matlab m-files for all the functions and scripts discussed in this course can be downloaded from the course website.

## Slide 17
Note: there are a number of forms used for the Fourier
Transforms. Some of these are different from the one we
will use here exclusively:
Some other examples:
(often seen in the
math literature)
Notes: Our definition of the Fourier transform and its inverse are not the only ones seen in the literature. Here, and on the next slide are some examples of different definitions.

## Slide 18
(often seen in the
EE literature)
All of these forms are acceptable. In fact we could write
Fourier transform pairs in general as:
as long as
Notes: All of these definitions are acceptable but we will only use our original choice in all subsequent work to be consistent. Choosing these other definitions will lead to differences in amplitude and/or phase from our definition so you must be careful when comparing our results with others that may appear in the literature.

## Slide 19
References
Sneddon, I.N., Fourier Transforms, McGraw-Hill, New York,
1951.
Bracewell, R.N., The Fourier Transform and its Applications,
3rd Ed. McGraw-Hill, New York, 2000.
Notes: Here are several basic references on Fourier transforms. There are many more.