## Slide 1
Introduction to NDE Ultrasonic
Systems
Notes: This set of slides is a general introduction to the basic elements of an ultrasonic nondestructive evaluation (NDE) system designed to examine materials or structures for flaws.

## Slide 2
Learning Objectives
familiarity with:
ultrasonic system components
ultrasonic NDE terminology
basic information on transducers
construction
types
common ultrasonic displays
Notes: We will examine the components that make up an ultrasonic NDE system and introduce you to some of the terminology that is commonly used in the ultrasonic NDE field.
Ultrasonic transducers are an important part of any system as they are used both to generate the ultrasound and to receive it. We will discuss some of the different types of transducers commonly used and how they are constructed.
There are various ways in which the signals received in an ultrasonic test are presented to the user so we will also describe some of the commonly used displays.

## Slide 3
flaw
cabling
cabling
Pulser
Receiver
flaw signal
Ultrasonic System Components
vi(t)
vR(t)
input voltage
output voltage
transducer
(transmitter)
transducer
(receiving)
component being
inspected
an ultrasonic NDE immersion flaw inspection
Notes: Here is a diagram of a basic ultrasonic NDE immersion system that uses two transducers placed in a water bath (one acts as a sound generator and one acts as a receiver) to examine a flawed component that is also in the water bath.
The pulser part of the system generates very short, repetitive voltage pulses (only one is shown) that travel through a cable and excite the sending transducer which converts those electrical pulses into pulses of sound that travel through the water and pass into the component.
These incident waves in the component can interact with any flaws present. The flaws, in turn, generate additional scattered ultrasound waves which in this setup are picked up by a separate transducer and converted back to voltage pulses that travel over a cable to a receiver where they are amplified and displayed as a voltage versus time signal on an oscilloscope.

## Slide 4
MODEL 5072PR PULSER-RECEIVER
Pulser/Receiver
Notes: This is a picture of a very basic pulser/receiver used in a lab setting. We will examine the pulser/receiver in more detail later so here we will just note that the left hand side of the instrument contains the controls for the pulser which include the pulse repetition frequency and energy and damping settings while the receiver side of the instrument contains a gain setting and filtering options. The back of the instrument (not shown) contain connections for feeding the received signals to an oscilloscope display.

## Slide 5
Pulse-echo
Pitch-catch
Common ultrasonic testing configurations
Notes: The pulser/receiver can either be used in a pulse-echo setting where a single transducer acts as both the transmitter and receiving transducer or in a pitch-catch setting where there are separate sending and receiving transducers.

## Slide 6
Through- transmission
Common ultrasonic testing configurations
Notes: It is also possible to use two transducers that are directly opposite to each other in a through-transmission setup. In this case note that even with a flaw absent, the receiving transducer will still be receiving a signal so that the flaw in this case acts as a modifier of this signal that passes through the component. In the pitch-catch case, in contrast, if the flaw is absent no signal will be received (at least from interactions with the flaw).

## Slide 7
Digital Oscilloscope
Notes: This is a picture of a typical digital oscilloscope which is used to digitally sample the received voltage versus time signals and display them on the oscilloscope screen. Since the signals are stored digitally, they can also be easily transferred to a computer for further processing.

## Slide 8
Ultrasonic Signals
Ultrasonic A-Scan
Notes: A voltage versus time trace on an oscilloscope display is also called an A-scan. It is the most commonly used type of display in ultrasonic tests.
Here we see a transducer in contact with the surface of a part receiving signals from a flaw in the component as well as a reflected signal from the back surface in a pulse-echo setup. On the oscilloscope screen we see the received signals as well as a large negative pulse at the start of the display called the main bang signal. This is a portion of the driving pulse that leaks through the pulser-receiver circuits to receiving side. This large signal can hide any flaw signals that may be present that are very close to the transducer so that there is an early part of the display, called a “dead zone” , where evaluation of the received signals may not be possible.

## Slide 9
Portable Flaw Detectors
Notes: Ultrasonic NDE tests are commonly done in the field, where a portable instrument is used. This instrument integrates the pulser/receiver and display into a single unit and provides the capability to perform various types of evaluations via the instrument settings.

## Slide 10
case
electrical network
backing
crystal
wear plate
connector
electrodes
Ultrasonic System Components
Contact Transducer
Notes: Here we see some typical ultrasonic transducers that are used in contact testing where the transducers are placed on the surface of a part and where a thin layer of water, oil, glycerin, or a commercially available couplant (not shown) is used to efficient couple the transducer to the part.
The transducer consists of a piezoelectric crystal which is plated on both faces. Acting as a transmitter, this crystal converts the electrical pulses acting on its faces into mechanical motion of the crystal, which then propagate through a wear plate and the couplant into the part as a traveling wave. The wear plate serves to protect the crystal from damage as the transducer is moved on the surface. A backing made of a highly attenuating material (such as epoxy loaded with small tungsten particles) dampens any waves traveling up into the transducer, thus preventing significant “ringing” of the crystal and generates a short pulse of ultrasound. There also may be some electrical components within the transducer casing that are used to adjust its electrical characteristics and “tune” the transducer output.
The same transducer can be used as a receiver which converts the motion of waves received into electrical signals.

## Slide 11
search tube
quarter wave plate
Ultrasonic System Components
Immersion Transducer
Notes: Here we see some transducers that are used in an immersion setup.
The basic construction of an immersion transducer is very similar to a contact transducer except the transducer usually has a UHF type of connector that is screwed into a search tube in a scanning apparatus. Also, the wear plate is replaced by a so-called quarter wave plate which is specifically designed to efficiently couple the crystal to a water bath. Unlike the contact transducer, where it is difficult to maintain a constant coupling to the part as the transducer is moved around, in an immersion setup the water provides a constant coupling medium regardless of the motion of the transducer. Thus, immersion tests are more reproducible than contact tests.

## Slide 12
Immersion
scanning
system
Notes: Here is a basic lab-top immersion setup where a transducer is connected to a scanning apparatus that allows the motion of the transducer to be controlled in a water tank. Much more sophisticated scanners can control both the angles and motions of the transducer.

## Slide 13
Ultrasonic System Components
Focused Transducer
acoustic lens
Notes: A planar (unfocused) immersion transducer generates a broad beam of ultrasound. However, if an acoustic lens is placed in front of the crystal, the sound can be focused into a part at a specific depth.

## Slide 14
Ultrasonic B-Scan
x
back surface
z
B-scan display
20 MHz B-Scan of a composite,
showing laminates
Notes: In addition to an A-scan, a B-scan is a commonly used type of display. Here, the x-location of the transducer is recorded as a transducer is moved along a line of the surface and the time of a received signal from a reflector is converted into the corresponding depth, z. A plot of x versus z on a display is called a B-scan. The picture on the right, for example, is a high frequency B-scan of a composite laminate where we can see the laminates clearly.

## Slide 15
Ultrasonic C-Scan
transducer scan pattern
x
y
z
C-scan image of chem-milled panels at
horizontal lap splice / vertical butt splice
on a 727 fuselage
Notes: In a C-scan a transducer is scanned in two dimensions on a surface and the (x, y) location is recorded as well as the response from any subsurface reflectors, which is usually color coded as a function of the amplitude of the response. An example C-scan is shown of the C-scan of a lap splice for a Boeing 727 fuselage.

## Slide 16
Ultrasonic System Components
References
Schmerr, L.W. and S.J. Song, Ultrasonic Nondestructive
Evaluation Systems – Models and Measurements, Springer,
2007.
Schmerr, L.W., Fundamentals of Ultrasonic Nondestructive Evaluation – A
Modeling Approach, 2nd Ed., Springer, 2016.
Schmerr, L.W., Fundamentals of Ultrasonic Phased Arrays, Springer, 2015.
Krautkramer, J. and H. Krautkramer, Ultrasonic Testing of
Materials, Springer Verlag, 1990.
Blitz, J. and G. Simpson, Ultrasonic Methods of
Non-destructive Testing, Chapman & Hall, London,UK,
1996.
Notes: Here are the three ultrasonic NDE books published by the author (L. Schmerr) of these notes, as well as several other commonly used general references.