## Slide 1
Cabling - Models and Measurements
Notes: Cabling can play a role in the measurement process for an ultrasonic NDE system so we will discuss the effects of cabling here in some detail.

## Slide 2
Learning Objectives
Cable transfer matrix model
Equivalent circuit model
Effects of termination conditions
Measurement of transfer matrix elements
Reciprocity check
Notes: We will model a cable as a two port system and show the effects of having different termination conditions. We will describe how we can measure the transfer elements of the cable directly and check to see if it acts as a reciprocal system.

## Slide 3
V1
I1
V2
I2
Cable – transfer matrix model
[T]
V1
I1
V2
I2
Transfer matrix from a simple 1-D transmission line model
Notes: At a fundamental level we can consider modeling the propagating electric and magnetic fields in a coaxial cable. However, we can use a simpler transmission line model where we consider the cable as a two port system that relates the voltages and currents at its ends though a 2x2 transfer matrix of components that depends on an electrical impedance of the cable, its length, and the wavenumber (or frequency).

## Slide 4
V1
I1
V2
I2
l
l = length of cable
kc = w/c = wave number (c = wave speed)
= electrical impedance
For an ideal coaxial cable:
b
a
e … permittivity
m … permeability
Cable – transfer matrix model
Notes: For an ideal circular coaxial cable we can write down an explicit expression for the impedance in terms of the radii of the inner and outer conductors and the electromagnetic properties of the material between those conductors.

## Slide 5
Equivalent circuit model of the transmission line model
I1
V1
I2
V2
-iZ0 tan(kl/2)
e
-iZ0 tan(kl/2)
e
-iZ0 / sin(kl/2)
e
Cable – transfer matrix model
Other equivalent circuits are also used
Notes: We can also represent the cable as two port system involving equivalent circuit components. The “T” type of equivalent circuit shown above is a popular choice but there are other equivalent circuits that can be used as well.

## Slide 6
[T]
V1
I1
Z2
Effects of different termination conditions
Cable – transfer matrix model
V1
I1
e
Z1
e
Notes: If we consider a cable that is terminated at one end, then we can characterize the terminated cable as an equivalent impedance, and by examining the ratio of the input voltage to the input current we can evaluate that equivalent impedance. Show are three cases where the cable is terminated by an impedance equal to the characteristic impedance of the cable and where the termination is either open circuit or closed circuit.

## Slide 7
Cable impedance measurements under different terminations
Cable – transfer matrix model
Notes: Here are some actual measurements of the magnitude and phase of the equivalent impedance of a 50 ohm coaxial cable when it is terminated by a 50 ohm resistor or is under open or closed circuit conditions. The equivalent impedance is measured by taking the ratio of the measured input voltage to the measured input current.

## Slide 8
[T]
V1
I1
Measuring the transfer matrix components
[T]
V1
I1
Is
V
T11 = V1 / V
T21 = I1 / V
T12 = V1 / Is
T22 = I1 / Is
open circuit
short circuit
Cable – transfer matrix model
Notes: By making various voltage and current measurements under different termination conditions we can also measure the individual elements of the 2x2 transfer matrix that characterizes the cable as a two port system.

## Slide 9
Measured cable transfer components
Cable – transfer matrix model
Notes: Here are the results of those measurements. It can be seen that the magnitudes of the T11 and T22 components have a cosine like behavior while the magnitudes of the T12 and T21 have a sine like behavior, consistent with the transmission line model. The phases also are consistent with that model

## Slide 10
Reciprocity Check
Cable – transfer matrix model
ideal reciprocal system
(amplitude)
ideal reciprocal
system (phase)
Notes: If we evaluate the determinant of the 2x2 transfer matrix we do get nearly a real value of unity for all frequencies, demonstrating the cable indeed is reciprocal.

## Slide 11
Homework problems
3.1, 3.2
Notes: Homework problems from Chapter 3.

## Slide 12
Staelin, D.H., Morgenthaler, A.W., and J. Kong,
Electromagnetic Waves,Prentice-Hall, 1994.
Balanis, C.A., Advanced Engineering Electromagnetics,
John Wiley, 1989.
Bladel, J.V., Electromagnetic Fields, Hemisphere Publuishing
Co., 1985.
Pozar, D.M., Microwave Engineering, John Wiley, 1998.
Karmel, P.R., Colef, G.D., and R.L. Camisa, Introduction to
Electromagnetic and Microwave Engineering, John Wiley, 1998.
Cable – transfer matrix model
Notes: Cables are discussed in a variety of EE texts. Here are some representative ones.