## Slide 1
Ultrasonic System Measurement Model - I
Notes: We now want to combine all the elements of an ultrasonic flaw measurement together and generate a complete ultrasonic measurement model.

## Slide 2
Learning Objectives
Use of a simple fluid model to obtain a complete
measurement model
extension to a fluid-solid model
Notes: Here, we will use a simple fluid model to develop the ultrasonic measurement model but we will show that this simple model captures all the elements of an immersion NDE measurement.

## Slide 3
flaw
pulser
Vi(w)
z i
transducer
Sr
transducer
receiver
eso
z s
eio
St
Ultrasonic System Model – Single Fluid
Notes: Here is the configuration we will examine: two transducers examining a scatterer (i.e. “flaw”) in a fluid in a pitch-catch measurement setup.

## Slide 4
pulser
Vi(w)
v0(w)
piston
transducer
Consider the sound generation process:
St … area of the transmitting transducer
Ultrasonic System Model – Single Fluid
transfer function for pulser, cabling, transducer
Notes: We have shown previously that we can write the compressive force generated by the sending piston transducer as a sound generation transfer function multiplied by the Thevenin equivalent voltage put out by the pulser. We saw how we could measure all the elements contained in the sound generation function (pulser impedance, cabling, transducer impedance and sensitivity).

## Slide 5
flaw
pulser
Vi(w)
Pinc(w)
z i
transducer
eio
In the paraxial approximation, the beam produces a quasi-plane wave incident on the flaw given by:
xs
St
Ultrasonic System Model – Single Fluid
Notes: In the paraxial approximation the beam of sound produced by the transducer generates a quasi-plane wave incident on the flaw.

## Slide 6
Ultrasonic System Model – Single Fluid
Pinc(w)
eio
xs
Note: since
ri
St
zi
diffraction correction
(deviation from a plane wave)
Notes: This quasi plane wave can be written in terms of plane wave traveling to the flaw multiplied by a diffraction correction terms. If we use, for example, a Rayleigh Sommerfeld representation of the beam then we can write that diffraction correction explicitly.

## Slide 7
St
ri
To calculate the diffraction coefficient, however, we can use a multi-Gaussian beam model:
a
Notes: In practice, however, it is easier to use a multi-Gaussian beam model to write the diffraction correction in more explicit terms.

## Slide 8
flaw
rs
For the waves scattered from the flaw
x
Pinc
eio
es
in the far field
Ultrasonic System Model – Single Fluid
plane wave far-field scattering amplitude
Notes: The scattered wave from the incident quasi-plane wave can be written in terms of the far field scattering amplitude.

## Slide 9
flaw
rs
Sr
receiving
transducer
Now, consider the scattered waves at the receiving
transducer:
es
Blocked force:
eio
Ultrasonic System Model – Single Fluid
Notes: The scattered waves ( which are spreading spherical waves) arrive at the receiving transducer. If we assume plane wave interactions at the receiving transducer, we can just double the scattered pressure field and compute the blocked force at the receiving transducer.

## Slide 10
flaw
eso
eio
rs
In the paraxial approximation
(on reception)
zs
Ultrasonic System Model – Single Fluid
and we have
receiving
transducer
Notes: Normally, the transducer is far enough from the flaw that the waves arriving are all traveling in essentially a single scattered direction so that the variation of the scattering amplitude over the face of the transducer can be ignored and the scattering amplitude taken out of the integral over the receiving transducer face.

## Slide 11
receiving
transducer
Ultrasonic System Model – Single Fluid
FB
blocked force
rs
If receiving transducer were acting as a transmitter it would
produce a pressure field given by
zs
Notes: Thus, the blocked force can be written as a product of the incident amplitude on the flaw times the far field scattering amplitude times a term which is again related to the diffraction correction for the receiving transducer when it is acting as a transmitter.

## Slide 12
receiving
transducer
Ultrasonic System Model – Single Fluid
FB
rs
zs
Thus,
and
becomes
Notes: Collecting all these results we have an explicit expression for the blocked force.

## Slide 13
receiving
transducer
Ultrasonic System Model – Single Fluid
FB
rs
zs
Again, we could compute this diffraction correction with
our multi-Gaussian beam model
a
Notes: Again it is easier to compute this diffraction correction with a multi-Gaussian beam model.

## Slide 14
transducer
receiver
Ultrasonic System Model – Single Fluid
FB
Sound reception process:
transfer function for transducer, cabling, receiver
Notes: We can relate the received voltage to the blocked force through a sound reception transfer function which accounts for the transducer, cabling and receiver.

## Slide 15
Combining all these relations, we find:
flaw
pulser
Vi(w)
z i
transducer
Sr
transducer
receiver
eso
z s
eio
St
Ultrasonic System Model – Single Fluid
FB
Notes: We now have terms in our measurement model, as shown, which consider the processes between the pulser and the flaw, the flaw and the receiving transducer, and from the receiving transducer to the measured received voltage.

## Slide 16
flaw
pulser
z i
transducer
Sr
transducer
receiver
eso
z s
eio
Ultrasonic System Model – Single Fluid
St
Vi(w)
Notes: Combining all the terms we have an (almost) complete measurement model. We say almost complete since we still need to account for wave attenuation, which we will do later.

## Slide 17
flaw
pulser
z i
transducer
Sr
transducer
receiver
eso
z s
eio
Ultrasonic System Model – Single Fluid
St
Vi(w)
… system function (pulser/receiver, cabling, transducer)
Notes: There are, of course, many measurements needed to obtain all the terms in the measurement model. But this burden can be greatly reduced by replacing many terms in the model by a single system function which can be measured in a calibration setup as we have discussed.

## Slide 18
pulser, receiver
cabling,transducers
flaw
scattering
beam propagation and diffraction effects going from the transmitting
transducer to the flaw
and from the flaw to the receiving transducer
received
voltage
Ultrasonic System Model-Single Fluid
When attenuation of the fluid is important, we need to also
include terms
a(w) … frequency dependent attenuation
area of
transmitter
Notes: Thus, the measurement model is a product of the system function (which can be measured), a set of terms involving beam propagation and diffraction effects for the sending and receiving transducers (which can be modeled with beam models), and the far field scattering amplitude, which characterizes the flaw response. Finally, there is a remaining term involving a set of constants. To complete this model we can also include wave attenuation terms involving propagation from the sending transducer to the flaw and from the flaw to the receiving transducer.

## Slide 19
Ultrasonic System Model-Single Fluid
For pulse-echo
For pitch-catch
Notes: Here we write the pitch-catch measurement model results in a more abbreviated fashion that shows the fields for the sending and receiving transducers explicitly and how the model reduces to a simpler form for the case of a pulse-echo measurement (sending and receiving transducers are identical)

## Slide 20
Thompson-Gray Measurement Model – Fluid-Solid (Pulse-Echo)
SA
z1
z2
fluid
solid
flaw
P-wave or S-wave
Notes: We have used a simple fluid model, but if we examine, for example the pulse-echo response of flaw in an actual ultrasonic measurement, a measurement model will be almost identical in form except we can look at P-wave or S-wave responses and the scalar scattering amplitude is a particular component of the vector scattering amplitude that depends on the polarization of the incident waves. The remaining constant coefficient is also slightly different. This measurement model was first obtained in a very similar form by Bruce Thompson and Tim Gray so it is called the Thompson-Gray measurement model.

## Slide 21
… propagation
… attenuation
… plane wave transmission
coefficient (velocity/velocity)
C (w) … diffraction coefficient
Thompson-Gray Measurement Model – Fluid-Solid (Pulse-Echo)
Notes: Here is the Thompson-Gray model with the velocity field expressed more explicitly in terms of propagation, attenuation, transmission, and diffraction terms.