## Slide 1
Pulser Characteristics
- Models and Measurements
Notes: This section will examine the pulser section of the pulser/receiver

## Slide 2
Learning Objectives
characteristics of "spike" and square wave pulsers
measurement of pulser electrical properties
effects of pulser settings
Notes: We will examine two types of pulsers used in practice – a “spike” pulser and a square wave pulser and discuss the measurement of the electrical characteristics of both of these types of instruments. We will also show some examples of the effects of different pulser settings.

## Slide 3
Pulser - Panametrics 5052 PR
rep rate
energy
attenuation
h.p. filter
damping
1
2
3
4
low
high
gain (dB)
power
T/R
RCVR
1
2
20
40
out
0.5
1.0
2.0
pulse-echo setup
transducer
Notes: Here is a diagram of a very simple pulser/receiver used in lab settings shown with an attached transducer which is being used as both a sender and receiver of ultrasound. The pulser settings on this instrument consist of the repetition rate setting, an energy level setting, and a damping control.

## Slide 4
Pulser - simple model of components
V
c
R
c
storage capacitor
(energy setting)
C
0
R
d
to
display
Q
0
damping
setting
voltage
source
repetitive
switch
(rep rate)
attenuation, gain
and filter settings
transducer (actually
an equivalent network)
receiver part
Notes: Here is a very basic model of what is in the pulser/receiver. A voltage source charges up a capacitor which then is periodically discharged. The amount of capacitance, and hence the energy stored, is controlled by the energy setting. The rep rate controls this discharge rate by the rate at which a switch is closed. One may want to reduce the rep rate in cases where one is inspecting a very large range of depths in a component so that all the flaw or other responses being recorded have attenuated and so that there is no overlap between the driving pulses and the received signal pulses. The damping setting controls a variable resistor where a low setting is for a high resistance and vice-versa. Although the transducer is a complex electromechanical device, because it consists of a piezoelectric crystal which is plated on its faces, it acts to first order like a capacitor, which is how it is characterized here. Obviously, the actual pulser circuits are much more complex than those shown here.

## Slide 5
Vi(w)
Z i(w)
e
V1
Pulser – Thevenin equivalent circuit
Vi(w) … equivalent voltage source
Z i(w) … equivalent impedance
e
Notes: Even though the pulser circuits may be very complex, if we assume the pulser acts as a linear system then we can replace it by a simpler Thevenin equivalent circuit consisting of a voltage source and an impedance.

## Slide 6
Many modeling studies omit the impedance and take
the voltage source as a negative time-domain “spike”:
Four parameter
model:
Pulser Modeling
Notes: A pulser such as the Panametrics 5052PR puts out very short “spikes” of voltage so it is called a spike pulser. Some modeling studies try to simulate the pulser output with a simple model such as the four parameter model shown here.

## Slide 7
Pulser Modeling
Model pulse with t0 = 0.01, a1 = 0.2, a2 = 50, V0 = 200
0
0.1
0.2
0.3
0.4
0.5
-200
-160
-120
-80
-40
0
Voltage
time, msec
Notes: Here is an example of the voltage pulse generated with this four parameter model for a set of parameters that mimic roughly what one sees as the output of the pulser: a downward going spike of about .05 microseconds duration and about 200 volts in amplitude.

## Slide 8
Pulser Modeling
Frequency spectrum calculated with FFT looks like:
0
20
40
60
80
100
120
140
160
180
200
0
1
2
3
4
5
frequency, MHz
|V(f)|
Notes: If one computes the frequency components of this pulse with an FFT, here is what the magnitude of the frequency components look like.

## Slide 9
>> t = linspace(0, 0.5, 512);
>> V = pulserVT(200, .005, 0.2, 50,t);
>> plot(t, V)
>> plot(t, c_shift(V, 100))
>> V = pulserVT(200, .01, 0.2, 50,t);
>> plot(t, c_shift(V, 100))
>> xlabel(' Time (\musec)')
>> ylabel('Voltage')
>> dt = t(2) -t(1);
>> Vf=FourierT(V,dt);
>> f = linspace(0, 1/dt, 512);
>> plot(f, abs(Vf))
>> plot(f(1:100), abs(Vf(1:100)))
>> xlabel('frequency, Hz')
>> ylabel('|V(f)|')
function V = pulserVT(V0, t0, a1, a2, t)
t = t + eps*( t ==0);
Vinf = V0/(1-exp(-a1*t0));
V = -Vinf*(1- exp(-a1*t)).*(t <= t0) -V0*exp(-a2*(t -t0)).*(t > t0);
Notes: Here is the MATLAB code that generates these time domain and frequency domain results.

## Slide 10
Equivalent voltage spectrum in the frequency domain:
A plot of the magnitude of this spectrum at the parameter values
listed previous looks like the previous plot:
Notes: However, we do not have to use the FFT to compute the frequency spectrum in this case since we can analytically perform the Fourier transform, which is shown here

## Slide 11
Pulser Modeling
0
20
40
60
80
100
120
140
160
180
200
0
1
2
3
4
5
6
frequency, MHz
|V(f)|
Notes: Here is a plot of the analytical Fourier transform which is indistinguishable from the FFT result.

## Slide 12
>> vff = pulserV(200, .01, 0.2,50,f);
>> plot(f(1:100), abs(vff(1:100)))
>> xlabel('frequency, Hz')
>> ylabel('|V(f)|')
function Vout = pulserV(V0, t0, a1, a2, f)
f = f + eps*(f ==0);
Vinf = V0/(1 -exp(-a1*t0));
arg1 = a1 - i*2*pi*f;
arg2 = a2 - i*2*pi*f;
Vout = Vinf*(1-exp(-arg1*t0))./arg1 + Vinf*(1 -exp(i*2*pi*f*t0))./(i*2*pi*f) ...
-V0*exp(i*2*pi*f*t0)./arg2;
Notes: Here is the MATLAB code for plotting this analytical result.

## Slide 13
Pulser – Panametrics 5052 PR
Output pulse of Panametrics 5052PR, energy setting = 1,
damping setting = 5, open output terminal
Notes: Here is the actual open circuit voltage measured for the Panametrics 5052PR pulser/receiver at the indicated energy and damping settings. It looks very much like our simulated four parameter mode. This is also the time domain voltage source in the Thevenin equivalent model.

## Slide 14
Vi(w)
Z i(w)
e
V
Measurement of pulser properties
Vi(w)
Z i(w)
e
VL
open circuit
V = Vi(w)
RL
known load, RL
Pulser - Experimental
Notes: We get the Thevenin equivalent voltage source in a Thevenin equivalent model from an open circuit voltage measurement and we can likewise determine the Thevenin equivalent electrical impedance by measuring the voltage across a known load, such as a resistance, placed across the output.

## Slide 15
Pulser – Panametrics 5052PR
Damping 0
Damping 7
Spike Pulser
0
5
10
15
20
0
10
20
30
40
50
60
70
f (MHz)
Magnitude (V)
Energy 1
Energy 4
f (MHz)
Magnitude (V)
0
5
10
15
20
0
2
4
6
8
10
12
14
16
Energy 1
Energy 4
Notes: Here are some pulser voltage source measurements in the frequency domain taken at different energy and damping settings

## Slide 16
Pulser – Panametrics 5052 PR
Damping 0
Damping 7
Spike Pulser
0
5
10
15
20
0
100
200
300
400
500
600
700
800
900
f (MHz)
Magnitude ()
Energy 1
Energy 4
0
5
10
15
20
0
5
10
15
20
25
30
f (MHz)
Magnitude ()
Energy 1
Energy 4
Notes: Similarly, here are some impedance measurements taken at different energy and damping settings

## Slide 17
Pulser – Panametrics 5052 PR
0
1
2
3
4
5
0
100
200
300
400
500
600
700
800
850
f (MHz)
Magnitude ()
Energy 1
Magnitude ()
f (MHz)
0
5
10
15
20
0
200
400
600
800
1000
1200
1400
1600
1800
2000
Damping effect
Spike Pulser
Damping
+
0
7
Loading effect
50W
Z
cable and
transducer
Notes: To get a better picture of the effects of the damping setting on the measured Thevenin equivalent impedance, on the left we see the frequency curves over a wide range of the damping values (at an energy 1 setting).
On the right we show the impedance calculated by measuring the voltage, VL, across a 50 ohm resistor (solid line) at the output of the pulser as well as the impedance calculated when a cable and transducer are attached to the pulser output and both the current and voltage (Vm, Im) at the output are measured (dashed line). In principle, we should get the same result in either case but we do see some differences depending on the loading of the pulser at its output port. In both cases Vi is the Thevenin equivalent voltage source. Thus, it might be advisable to conduct these measurements under loading conditions actually present in an experiment.

## Slide 18
Square Wave Pulser -UTEX 340
Notes: Another type of pulser/receiver commonly used in practice contains a square wave pulser such as the UTEX 340. This instrument can be controlled from its front panel or from an equivalent computer interface, as shown. In this case there are three pulser controls consisting of the pulse repetition rate, the pulse width, and the pulse voltage. The gain and high and low pass filter controls are for the receiver section.

## Slide 19
Square Wave Pulser -UTEX 320
t (sec)
Amplitude (V)
f (MHz)
Amplitude (V)
Phase (deg)
Pulser source measured in open-circuit conditions
0
0.05
0.1
0.15
0.2
0.25
-200
-100
0
Notes: Here is the measured open circuit voltage and its frequency components, showing the square-like nature of the pulse in the time-domain

## Slide 20
Square Wave Pulser -UTEX 320
Pulse Width 0
Pulse Width 50
0
5
10
15
20
0
1
2
3
4
5
6
7
f (MHz)
Magnitude (V)
100 V
200 V
f (MHz)
Magnitude (V)
0
5
10
15
20
0
2
4
6
8
10
12
14
100 V
200 V
Square wave Pulser
Notes: This shows as the pulse width widens the frequency domain response does become less broad.

## Slide 21
Square Wave Pulser -UTEX 320
Pulse Width 0
Pulse Width 50
0
5
10
15
20
0
20
40
60
80
100
120
140
160
f (MHz)
Magnitude ()
100 V
200 V
Square wave Pulser
f (MHz)
Magnitude ()
0
5
10
15
20
0
5
10
15
20
25
30
35
40
100 V
200 V
Notes: Here are some equivalent impedance measurements under different pulse width and pulse amplitude settings.

## Slide 22
Square Wave Pulser -UTEX 320
f (MHz)
Magnitude ()
Voltage 100
Loading effect
Magnitude ()
f (MHz)
Pulse width effect
20
0
5
10
15
20
0
40
60
80
100
120
140
160
180
200
220
Square wave Pulser
Width
+
10
50
50W
Z
cable and
transducer
Notes: On the left we see the effects of changing pulse width settings on the equivalent impedance.
On the right we see the dependency of this impedance on the loading used at the output terminal. Again, the first case (solid line) is where a 50 ohm resistor is placed at the output and the voltage VL is measured, while the second case (dashed line) is where a cable and transducer are attached an the voltage and current are both measured.

## Slide 23
Homework problem
2.1
Notes: Homework problem from chapter 2.