## Slide 1
Transducer Characterization
Notes: Here we will examine the measurements needed to fully characterize the transducer(s) used in an ultrasonic measurement system.

## Slide 2
Learning Objectives
Pulse-echo experimental determination of impedance,
sensitivity of commercial transducers
Experimental determination of effective radii and
focal lengths of commercial transducers
Notes: The transducer parameters we will need to obtain experimentally are the transducer electrical impedance, the transducer sensitivity, and the effective radius and focal length (if focused) of the transducer.

## Slide 3
Transducer impedance, sensitivity
To fully characterize these generation and
reception transfer functions, we need to be
able to obtain the transducers input electrical
impedance and their sensitivities
Notes: Again, recall that the transducer electrical impedance and sensitivity are the two key parameters we need to find to determine the transducer’s role in the measurement process.

## Slide 4
Transducer impedance
Measurement of the transducer input impedance, Zin :
Pulser
Transducer
Current probe
V
I
water
Short cable
Notes: The electrical impedance is easy to obtain if we simply connect a transducer and short cable to the pulser and measure the voltage and current in the cable before any reflected signals are received. Taking the Fourier transform of those signals then we find the impedance directly.

## Slide 5
Transducer impedance
0
5
10
15
20
0
500
1000
1500
0
5
10
15
20
80
90
100
110
120
130
140
150
Frequency (MHz)
Phase (Deg)
Amplitude (W)
amplitude
phase
Example impedance measurements of two 5 MHz transducers:
transducer 1
Notes: We will show impedances measurements for two 5 MHz transducers. Here are the results for the first transducer.

## Slide 6
Transducer impedance
0
5
10
15
20
0
500
1000
1500
0
5
10
15
20
80
90
100
110
120
130
140
150
Frequency (MHz)
Phase (Deg)
Amplitude (W)
amplitude
phase
transducer 2
Notes: Here are the results for the second transducer.

## Slide 7
0
1
2
3
4
5
6
7
8
9
10
0
2
4
6
8
frequency
0
1
2
3
4
5
6
7
8
9
10
89
89.5
90
90.5
91
Transducer impedance
Impedance of a capacitor, Z =1/(-iwC)
magnitude
phase
Notes: Both of those transducers have a frequency dependent behavior that looks much like that of a capacitor. Here is a capacitor’s electrical impedance. This is not surprising since a transducer is a piezoelectric crystal that is plated on both faces.

## Slide 8
pulser/receiver
Transducer sensitivity
Pulse-echo measurement setup for determining sensitivity
solid
Notes: To determine the sensitivity we can use a pulse-echo immersion setup where we place the transducer in a water bath at normal incidence to the flat face of a solid block.

## Slide 9
Transducer sensitivity
solid
a
b
1. measure voltage and current when transducer is radiating into the fluid but before any reflected waves have arrived
Notes: First, we measure the voltage and current at the pulser/receiver before any reflected waves have arrived.

## Slide 10
Transducer sensitivity
2. do FFT of the measured voltage and current and relate to the voltage and current at the transducer by compensating for the cabling
Note: then the impedance of the transducer is just
Notes: We do the FFT of those voltage and current measurements and then, for a known cable, compensate for the cabling effects to obtain the corresponding voltage and current at the input port of the transducer. The ratio of this voltage and current again is just the electrical impedance, but here we do not have to use a short cable to obtain that impedance.

## Slide 11
Transducer sensitivity
3. measure voltage and current generated by the waves reflected from the surface of the block:
solid
a
b
Notes: Then we measure the received voltage and current at the pulser/receiver generated by the waves reflected from the block

## Slide 12
Transducer sensitivity
4. do FFT of the measured voltage and current and relate to the voltage and current at the transducer electrical port by compensating for the cabling
Notes: We then do an FFT of the measured reflected wave signals (called V2 , I2 here) and compensate for the cabling to get the voltage and current ( VT , IT ) at the transducer electrical port. Note the changes in sign consistent with the definitions of inputs and outputs and the assumed directions of the currents. Ideally, the diagonal elements of the cable would be identical and the cable would be completely reciprocal but we have allowed for small differences in the actual measurements.

## Slide 13
Transducer sensitivity
5. From these measurements and a knowledge of the acoustic/elastic transfer function for this setup we can obtain the sensitivity of the transducer from:
[Note: in all these division processes in the frequency domain, a Wiener filter is used to desensitize the process to noise]
Notes: One can show that the transducer sensitivity can be obtained from these measurements since we can model the acoustic/elastic transfer function for this setup and the acoustic radiation impedance of the transducer. We will not show those models here.

## Slide 14
Example sensitivity calculated for a
5MHz planar transducer:
Transducer sensitivity
0
0.05
0.1
0.15
0.2
0.25
0.3
0.35
Amplitude (V/N)
f (MHz)
0
5
10
15
20
-200
-150
-100
-50
0
50
100
Phase (deg)
Notes: Here is an example sensitivity calculated for a 5 MHz transducer.

## Slide 15
Transducer sensitivity
Cabling effects need to be accounted for in determining the sensitivity:
0
5
10
15
20
0
0.1
0.2
0.3
Amplitude (V/N)
0
5
10
15
20
-200
0
200
400
f (MHz)
Phase (deg)
Compensation
No compensation
f (MHz)
Notes: The compensations for cabling were important in these measurements as shown here, where the results when cabling effects are ignored are also shown. We could, of course, also eliminate those cabling effects by doing the measurements directly at the transducer’s electrical port, but because this is an immersion setup those measurements would have to be taken under the water and it is more practical to do the measurements at the pulser/receiver.

## Slide 16
Transducer Effective Parameters
Use of manufacturer specifications for parameters such as
transducer radius and focal length in transducer models
do not always lead to good agreement with experimental
measurements.
a
R0
Thus, we need to determine experimentally "effective" values
for these parameters.
Notes: A manufacturer will specify quantities such as transducer radius and focal length for a commercial transducer but those values may not be able to be used directly in models of ultrasonic measurements without causing errors, which can sometimes be large. Thus, we need to be able to determine experimentally these parameters, which we will call effective parameters.

## Slide 17
Transducer Effective Parameters
Effective radius for an unfocused (piston) immersion probe
sample
FFT
5 MHz
0
0.5
1
1.5
2
2.5
3
3.5
4
0
0.2
0.4
0.6
0.8
1
1.2
1.4
1.6
1.8
2
|p|
rcv0
z/N
plot 5 MHz component
versus z
z
find zmin
determine
aeff
a
Notes: We can measure the effective radius of an unfocused circular transducer by placing a small reflector in a immersion tank and moving that reflector at different distances along the central axis of the transducer. A frequency domain model of the on-axis pressure distribution on the axis of the transducer shows that in the so-called near field of the transducer there are series of peaks and nulls in the pressure at a given frequency. Those pressure variations will cause the measured frequency components of the received voltage from the reflector also to vary. Thus, if one measures the time domain waveform received from the reflector at a given distance and does an FFT to get the frequency components of the signal, then one can look at the response at a given frequency (say 5MHz as shown above). If this same 5 MHz response is obtained when the reflector is at different distances, one can locate the distance, zmin , where the 5 MHz response has its last zero value. Since this distance is related to the radius of the transducer and the wavelength (and hence, the frequency) one can use it to predict an effective radius value. In principle this effective radius should not depend on the choice of frequency but experiments show that there can be some variation with frequency so if that is the case then one usually uses an average effective radius value calculated at different frequencies over the bandwidth of the transducer.

## Slide 18
Effective radius and focal length for a spherically focused probe
0
0.5
1
1.5
2
2.5
3
3.5
4
0
2
4
6
8
10
12
|p|
rcv0
zmin
zmax
where x is the root of:
Transducer Effective Parameters
Notes: For a spherically focused transducer we must determine both an effective radius and an effective focal length so we will need at least two measured values. Again, we can use the same setup just described and obtain the magnitude of the response at a given frequency for a reflector at different distances along the transducer central axis. A model of a spherically focused circular transducer shows that the transducer has an on-axis pressure response as shown above. Here, one measures the location of the large peak of the response (which is called the true focus) and the same last on-axis null. The model then also shows that we can obtain the effective radius and focal length values from those two measurements. The null location is usually easy to measure but the exact location peak is a bit more difficult to obtain so one can use a number of possible peak locations to try to determine those effective values that best fit the on-axis response at other points as well.

## Slide 19
Effective transducer parameters determined experimentally:
Probes | Manufacturers Specs focal length radius (cm) (cm) | Estimated Parameters focal length radius (cm) (cm) | Center Frequency (MHz)
A | 7.62 0.476 | 13.47 0.451 | 10
B | 7.62 0.635 | 20.74 0.556 | 5
C | 7.62 0.476 | 7.45 0.469 | 15
These values should be independent of frequency but in
practice they do vary somewhat with the frequency component
used in their determination.
Transducer Effective Parameters
Notes: Here are some example calculations of effective transducer parameters. These values were obtained by measuring the on-axis response at the center frequency of the transducer. Other frequencies may give somewhat different values. One can see that in some cases the results are quite different from the manufacturer’s specifications.

## Slide 20
Homework problem
6.1
Notes: Homework problem from Chapter 6.