## Slide 1
Ultrasonic System Overview
Notes: These slides are essentially Chapter 1 in the book: Schmerr, L.W. and S.J. Song, Fundamentals of Ultrasonic Evaluation Systems – Models and Measurements.

## Slide 2
Learning Objectives
Overview of the UT system components that will be discussed
Notes: In subsequent presentations we are going to discuss all the elements that make up an ultrasonic NDE inspection system. This set of slides will walk through an entire system and give a preview of what is to come.

## Slide 3
flaw
pulser/
receiver
cabling
oscilloscope display
receiving transducer
transmitting
transducer
The ultrasonic system components
Notes: Here are all the components we will discuss in detail: the pulser/receiver, the cabling, the transducers, and the oscilloscope display. An immersion setup is shown only to illustrate these components.

## Slide 4
-2.5
-2
-1.5
-1
-0.5
0
0.5
1
1.5
2
2.5
-1
-0.6
-0.2
0.2
0.6
1
time,
m
seconds
0
1
2
3
4
5
6
7
8
9
10
0
0.1
0.2
0.3
0.4
frequency, MHz
(a)
(b)
amplitude, volts
amplitude, volts/MHz
voltage versus time pulse, v(t)
magnitude of the frequency spectrum of the pulse, |V(f)|
(only positive frequencies shown)
Fourier
Transforms
f = frequency
(in cycles/sec or Hz)
w = 2pf
=frequency
(in rad/sec)
Notes: Although ultrasonic systems generate and receive pulses (short signals as a function of time) we can use a Fourier transform to generate the frequency components contained in the pulses and plot the so-called spectrum of the pulses as a function of frequency. Shown is an example of a short signal and its spectrum as might be seen in an ultrasonic NDE test. If we know the spectrum of the signal we can also use an inverse Fourier transform to recover the time signal.
Frequency is usually measured either in terms of Hertz (cycles/sec) or in terms of omega (radians/sec). In NDE tests, the very high frequencies involved are typically measured in millions of Hertz, or MHz, which stands for megaHertz

## Slide 5
pulser
The ultrasonic pulser and its equivalent circuit
voltage source
electrical impedance
pulser
voltage, current
cable
voltage pulse
v(t)
(also pulse of current, i(t))
V(w)
I(w)
Notes: The pulser section of a pulser/receiver puts out repetitive short pulses of voltage and current as a function of time (only one is shown). If we compute the frequency components of these signals we can show that we can model how the pulser generates these outputs in the frequency domain by replacing the complex circuits contained in the pulser with a model that represents the pulser simply as a voltage source and an electrical impedance.

## Slide 6
transmitting cabling and its equivalent transfer matrix
Notes: Normally we think of a cable as simply transferring signals from one location to another. However, at the MHz frequencies seen in ultrasonic NDE tests we will see this is not in general true unless the cable is short (roughly about a meter or less). In the frequency domain the cable can be modeled as a two port system where input voltage and current are related to the output voltage and current through a 2x2 transfer matrix. This transfer matrix can be determined though a series of electrical measurements.

## Slide 7
(a)
[TA]
(b)
A
a sending transducer and its equivalent transfer matrix
voltage, current
force, velocity
Notes: A transmitting transducer can also be modeled as a two port system where the voltage and current inputs are converted into mechanical outputs of force and velocity. Again, a 2x2 transfer matrix will characterize this two port system.

## Slide 8
Vin
Iin
(a)
(b)
electrical impedance
sensitivity
acoustic radiation
impedance
a sending transducer modeled as an impedance and sensitivity
Notes: Unlike a cable, it is difficult to measure directly the transfer matrix components of a transducer with purely electrical measurements. However, because such a transducer is always used when the acoustic output side of the transducer is in contact with some medium (fluid or solid), the output force and velocity are always related to each other through an acoustic radiation impedance, i.e. the output of the transducer is mechanically terminated. Under these conditions we can replace the two port system with a transducer model consisting of an electrical impedance and a sensitivity, both of which can be determined with electrical measurements.

## Slide 9
output
force
(a)
(b)
pulser
A
(c)
putting all the sound generation components together
sound generation transfer function
Notes: Models of all the elements contained in the sound generation process can be combined as shown. These elements can all be lumped into a single complex sound generation transfer function that relates the input voltage spectrum of the pulser to an acoustical output of the sending transducer such as the force. Thus, we can reduce the entire sound generation process into a model of single input-single output system characterized by the sound generation transfer function.

## Slide 10
(a)
B
incident and
scattered waves
Modeling the acoustic sources acting on the receiving transducer
blocked force
acoustic radiation impedance
Notes: When a transducer is used a receiver, the waves incident on and scattered from the transducer act as acoustic driving terms as shown in (a). These waves can be modeled as a acoustic source, called the blocked force (which will be defined later) acting on the transducer face and the acoustic radiation impedance of the receiving transducer.

## Slide 11
voltage
impedance
the receiving transducer also can be modeled by an impedance and sensitivity and these combined with the acoustic sources to yield a simple source and impedance model:
Notes: The 2x2 transfer matrix of the receiving transducer can also be replaced by an electrical impedance and sensitivity and these elements can be combined with the acoustic driving terms to model the receiving transducer and its acoustic sources by simply an equivalent voltage source and an electrical impedance.

## Slide 12
receiving cabling and its equivalent transfer matrix
Notes: A cable between the receiving transducer and the receiver section of the pulser/receiver can again be modeled as a two port system characterized by a 2x2 transfer matrix.

## Slide 13
model of the receiver as an electrical impedance and gain factor
impedance
gain factor
receiver
output voltage
Notes: The receiver section of a pulser/receiver acts to amplify the received signals and often provide low and high pass filters that modify the frequency content of the signals. In quantitative NDE studies we normally do not want to filter out any of the available information so that we will model the receiver as only an electrical impedance and a frequency dependent gain factor. If one wishes to add filters to this model, that is easily accomplished.

## Slide 14
receiver
output
voltage
(a)
(b)
(c)
putting all the sound reception components together
Notes: One can combine the receiving transducer, cable, and receiver into a composite model where all the elements are known. Again, these elements can be combined into a receiving transfer function that relates the blocked force to the received voltage in a single input, single output system that characterizes the receiving part of the entire system.

## Slide 15
acoustic/elastic
processes
pulser
receiver
Modeling all the wave propagation and scattering components as a transfer function
acoustic/elastic transfer function
Notes: Since we have models of both the sound generation and sound reception parts of a measurement system, to complete the model for the entire system we need to characterize all the acoustic/elastic waves present between the transducers in terms of an acoustic/elastic transfer function that relates the force generated by the transmitting transducer to the blocked force acting on the receiving transducer.
Unlike the sending and receiving transfer functions, however, it is not possible to measure the acoustic/elastic transfer function since it due to complex 3-D wave interactions that are not directly accessible. Thus, one must model this transfer function with appropriate wave propagation and scattering models.

## Slide 16
pulser
receiver
T
R
flaw
Modeling an entire ultrasonic measurement system
sound generation
transfer function
sound reception
transfer function
Notes: If we combine together all the elements we have described we now have a complete model of an ultrasonic NDE measurement system.

## Slide 17
system function
acoustic/elastic transfer function
output voltage
In terms of transfer functions the ultrasonic system looks like:
We can also combine the sound generation and reception transfer functions with the voltage source of the pulser tin a single function called the system function:
Notes: A complete ultrasonic system model can thus be characterized in terms of the three transfer functions we have discussed. However, we can combine the sound generation and sound reception transfer functions together with the voltage source of the pulser into a single system function that we will show we can measure directly without knowing all the individual components that go into that system function. Thus, ultimately we can simply characterize a complete ultrasonic system as a single input, single output system where if we measure the system function and model the acoustic/elastic transfer function, we can predict the spectrum of the received voltage signals of the entire measurement system. An inverse Fourier transform of this spectrum then will produce the measure signals as a function of time.